## Uncertainty
Boonton RTP5000 Real-Time USB Peak Power Sensor Measurement Uncertainty Calculator v0.3
Power Sensor
Select USB Power Sensor | RTP5006 | Measurement Range (dBm) |  | Frequency Range (GHz)
Sensor VSWR | 1.25 | -60 | 20 | 0.05 | 6
VSWR @ cal frequency | 1.25
Device Under Test (DUT)
Measured Power Level | 0 | dBm | 1 | 2
Measured Frequency | 1 | GHz | 2
VSWR | 1
|  |  |  | 1.4 | 2
Measurement Settings |  |  |  | 1.4
Measurement Type (PULSE or CW) | CW
Calibration Temperature (C) | 22 | Actual detector internal temperature
Measurement Temperature (C) | 22 | Actual detector internal temperature
Sources of Uncertainties
Power Sensor | Estimate (in %) | Probability Distribution |  | Divisor | Standard Uncertainty (in %)
Instrument Uncertainty | 0.2 | Normal |  | 2 | 0.1
Calibrator Level Uncertainty | 0.9 | Rectangular |  | 1.7321 | 0.51962
Calibrator Mismatch Uncertainty | 0.54201 | U-shaped |  | 1.4142 | 0.38326
Sensor Shaping Error | 0.5 | Rectangular |  | 1.7321 | 0.28868
Temperature Drift | 0.93 | Rectangular |  | 1.7321 | 0.53694
Noise and Drift | 0.0001 | Normal |  | 2 | 5e-05
Cal Factor Uncertainty | 1.4 | Normal |  | 1 | 1.4
Device-Under-Test
Mismatch between Sensor and DUT | 0 | U-shaped |  | 1.4142 | 0
Measurement Uncertainty
Combined Uncertainty (%) | Root_Sum_Squared |  |  |  |  | 1.6609
Expanded Uncertainty (%) | Coverage Factor (K) |  | 2 | Normal | ± | 3.3218
Expanded Uncertainty (dB) | Coverage Factor (K) |  | 2 | -0.14671 | to | 0.14192

## Sensor Data (hidden)
| RTP5006 |  | RTP5318 |  | RTP5518 |  | RTP5340 |  | RTP5540 |  | RTP5008 ≤6 GHz |  | RTP5008 >6 GHz |  | Available Sensors
Freq Range (GHz) | [14 empty] | RTP5006
Low |  | 0.05 |  | 0.05 |  | 0.05 |  | 0.05 |  | 0.05 |  | 0.05 |  | 6 | RTP5318
High |  | 6 |  | 18 |  | 18 |  | 40 |  | 40 |  | 6 |  | 8 | RTP5518
Measurement Range (dBm) | [14 empty] | RTP5340
Pulse | [14 empty] | RTP5540
Low |  | -50 |  | -24 |  | -40 |  | -24 |  | -40 |  | -50 |  | -43 | RTP5008 ≤6 GHz
High |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 | RTP5008 >6 GHz
CW
Low |  | -60 |  | -34 |  | -50 |  | -34 |  | -50 |  | -60 |  | -53
High |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20
Sensor VSWR
| 0.05 | 1.25 | 0.05 | 1.15 | 0.05 | 1.15 | 0.05 | 1.25 | 0.05 | 1.25 | 0.05 | 1.2 | 0.05 | 1.2
| 6 | 1.25 | 2 | 1.15 | 2 | 1.15 | 4 | 1.25 | 4 | 1.25 | 6 | 1.2 | 6 | 1.2
|  |  | 16 | 1.28 | 6 | 1.2 | 38 | 1.65 | 38 | 1.65 |  |  | 8 | 1.25
|  |  | 18 | 1.34 | 16 | 1.28 | 40 | 2 | 40 | 2
[5 empty] | 18 | 1.34
Drift and Noise (nW)
PULSE | PULSE | 10 | PULSE | 4000 | PULSE | 50 | PULSE | 4000 | PULSE | 50 | PULSE | 10 | PULSE | 40
CW | CW | 1 | CW | 400 | CW | 5 | CW | 400 | CW | 5 | CW | 1 | CW | 4
Shaping Error (%) | 0.5 | 2.3 | 0.5 | 2.3 | 0.5 | 4 | 0.5 | 2.3 | 0.5 | 4.7 | 0.5 | 2.3 | 0.5 | 3.5
Cal Factor Uncertainty (%RSS) | Tegam ↓
Freq (GHz)
| 0.5 | 1.4 | 0.5 | 1.4 | 0.5 | 1 | 0.5 | 1.7 | 0.5 | 1.5 | 0.5 | 1.4 | 0.5 | 1.4
| 1 | 1.4 | 1 | 1.4 | 1 | 1.1 | 1 | 1.3 | 1 | 1.1 | 1 | 1.4 | 1 | 1.4
| 2 | 1.4 | 2 | 1.4 | 2 | 1.5 | 2 | 1.3 | 2 | 1.2 | 2 | 1.4 | 2 | 1.4
| 3 | 1.5 | 3 | 1.4 | 3 | 1.5 | 3 | 1.5 | 3 | 1.3 | 3 | 1.5 | 3 | 1.5
| 4 | 1.5 | 4 | 1.5 | 4 | 1.3 | 4 | 1.7 | 4 | 1.5 | 4 | 1.5 | 4 | 1.5
| 5 | 1.5 | 5 | 1.5 | 5 | 1.7 | 5 | 1.8 | 5 | 1.5 | 5 | 1.5 | 5 | 1.5
| 6 | 1.5 | 6 | 1.5 | 6 | 1.5 | 6 | 2 | 6 | 1.6 | 6 | 1.5 | 6 | 1.5
|  |  | 7 | 1.5 | 7 | 1.1 | 7 | 2.3 | 7 | 1.9 |  |  | 7 | 1.7
|  |  | 8 | 1.6 | 8 | 1.1 | 8 | 2.6 | 8 | 2.2 |  |  | 8 | 1.7
|  |  | 9 | 1.5 | 9 | 1.7 | 9 | 4.4 | 9 | 3.8
|  |  | 10 | 1.5 | 10 | 2.2 | 10 | 4.7 | 10 | 4.1
|  |  | 11 | 1.6 | 11 | 2.5 | 11 | 5 | 11 | 4.4
|  |  | 12 | 1.7 | 12 | 2.4 | 12 | 5 | 12 | 4.4
|  |  | 13 | 1.6 | 13 | 2.5 | 13 | 5 | 13 | 4.4
|  |  | 14 | 1.5 | 14 | 2.7 | 14 | 5 | 14 | 4.5
|  |  | 15 | 1.6 | 15 | 2.5 | 15 | 4.6 | 15 | 4.4
|  |  | 16 | 1.7 | 16 | 3 | 16 | 4.1 | 16 | 4
|  |  | 17 | 1.6 | 17 | 2.5 | 17 | 3.7 | 17 | 3.7
|  |  | 18 | 1.6 | 18 | 2.4 | 18 | 3.5 | 18 | 3.7
[7 empty] | 19 | 3.6 | 19 | 4.2
[7 empty] | 20 | 3.7 | 20 | 4.4
[7 empty] | 21 | 4.2 | 21 | 4.6
[7 empty] | 22 | 4.6 | 22 | 4.8
[7 empty] | 23 | 4.8 | 23 | 5
[7 empty] | 24 | 4.8 | 24 | 5.2
[7 empty] | 25 | 4.8 | 25 | 4.9
[7 empty] | 26 | 5 | 26 | 5.2
[7 empty] | 26.5 | 5.1 | 26.5 | 5.2
[7 empty] | 27 | 5.6 | 27 | 5.8
[7 empty] | 28 | 5.1 | 28 | 5.3
[7 empty] | 29 | 4.8 | 29 | 5.1
[7 empty] | 30 | 4.7 | 30 | 5
[7 empty] | 31 | 4.9 | 31 | 5.1
[7 empty] | 32 | 5.3 | 32 | 5
[7 empty] | 33 | 5.6 | 33 | 5
[7 empty] | 34 | 5.1 | 34 | 4.7
[7 empty] | 35 | 5.3 | 35 | 4.6
[7 empty] | 36 | 4.7 | 36 | 4.5
[7 empty] | 37 | 4.3 | 37 | 4.6
[7 empty] | 38 | 4.7 | 38 | 5.7
[7 empty] | 39 | 5.4 | 39 | 7.5
[7 empty] | 40 | 5.6 | 40 | 7.9
Calibration Frequency (GHz) |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1
Power Level Standard UNC |  | 0.9 |  | 0.9 |  | 0.9 |  | 0.9 |  | 0.9 |  | 0.9 |  | 0.9
Calibrator Match |  | 1.05 |  | 1.05 |  | 1.05 |  | 1.05 |  | 1.05 |  | 1.05 |  | 1.05
When using the Tegam system, this match is better, giving us a better uncertainty at 0 dBm.
However, the relative uncertainties at different levels will be worse because we are not stepping power.
Neither of these effects are captured yet in our uncertainty values provided to customers.
1?

## Formulas (hidden)
B5
RTP5006 | 1.25
RTP5318 | 1.15
RTP5518 | 1.15
RTP5340 | 1.25
RTP5540 | 1.25
RTP5008 ≤6 GHz | 1.2
RTP5008 >6 GHz | 1.2
B6
RTP5006 | 1.25
RTP5318 | 1.15
RTP5518 | 1.15
RTP5340 | 1.25
RTP5540 | 1.25
RTP5008 ≤6 GHz | 1.2
RTP5008 >6 GHz | 1.2
B21
RTP5006 | 0.9
RTP5318 | 0.9
RTP5518 | 0.9
RTP5340 | 0.9
RTP5540 | 0.9
RTP5008 ≤6 GHz | 0.9
RTP5008 >6 GHz | 0.9
B22
RTP5006 | 0.54201
RTP5318 | 0.54201
RTP5518 | 0.54201
RTP5340 | 0.54201
RTP5540 | 0.54201
RTP5008 ≤6 GHz | 0.54201
RTP5008 >6 GHz | 0.54201
B23
RTP5006 | 0.5
RTP5318 | 0.5
RTP5518 | 0.5
RTP5340 | 0.5
RTP5540 | 0.5
RTP5008 ≤6 GHz | 0.5
RTP5008 >6 GHz | 0.5
B25
RTP5006 | 0.0001
RTP5318 | 0.04
RTP5518 | 0.0005
RTP5340 | 0.04
RTP5540 | 0.0005
RTP5008 ≤6 GHz | 0.0001
RTP5008 >6 GHz | 0.0004
B26
RTP5006 | 1.4
RTP5318 | 1.4
RTP5518 | 1.1
RTP5340 | 1.3
RTP5540 | 1.1
RTP5008 ≤6 GHz | 1.4
RTP5008 >6 GHz | 1.4
C5
RTP5006 | -60
RTP5318 | -34
RTP5518 | -50
RTP5340 | -34
RTP5540 | -50
RTP5008 ≤6 GHz | -60
RTP5008 >6 GHz | -53
D8
RTP5006 | 1
RTP5318 | 1
RTP5518 | 1
RTP5340 | 1
RTP5540 | 1
RTP5008 ≤6 GHz | 1
RTP5008 >6 GHz | 1
D9
RTP5006 | 2
RTP5318 | 2
RTP5518 | 2
RTP5340 | 2
RTP5540 | 2
RTP5008 ≤6 GHz | 2
RTP5008 >6 GHz | 2
E11
RTP5006 | 1.4
RTP5318 | 1.4
RTP5518 | 1.1
RTP5340 | 1.3
RTP5540 | 1.1
RTP5008 ≤6 GHz | 1.4
RTP5008 >6 GHz | 1.4
E12
RTP5006 | 1.4
RTP5318 | 1.4
RTP5518 | 1.5
RTP5340 | 1.3
RTP5540 | 1.2
RTP5008 ≤6 GHz | 1.4
RTP5008 >6 GHz | 1.4
E8
RTP5006 | 2
RTP5318 | 2
RTP5518 | 2
RTP5340 | 2
RTP5540 | 2
RTP5008 ≤6 GHz | 2
RTP5008 >6 GHz | 2
F11
RTP5006 | 2
RTP5318 | 2
RTP5518 | 2
RTP5340 | 2
RTP5540 | 2
RTP5008 ≤6 GHz | 2
RTP5008 >6 GHz | 2
F5
RTP5006 | 6
RTP5318 | 18
RTP5518 | 18
RTP5340 | 40
RTP5540 | 40
RTP5008 ≤6 GHz | 6
RTP5008 >6 GHz | 8